## Slide 1
Annual Internal Quality Audit - Closing Meeting
Lead Auditor : Simy Li
Lead Auditor Trainee : Paul Ip
Auditor : Kelvin Tsoi
Auditor Trainees : Brian Wong, Shannon Tsui (CAS),
Cammy Hung (Lab)
Date : March 17, 2011

## Slide 2: Observation 1 – from Paul Ip
Project 10CA39214 (WTDP Program)
Requirement:00-LC-S0258 issue 16 (Clause 9.4H-4)
NC: Datasheet mention that the fan shall be installed at the height of at least 2.1 m from ground to the lowest point of fan blade. But not found this verification result (2.1 m) and calibrated equipment in Datasheet.

## Slide 3: Observation 2 – from Paul Ip
Project 10CA52893 (CTDP Program)
Requirement: 00-CE-S0030 & 00-LC-S0258
NC:
CTDP DS was not associated to DAP database
Not using “strikethrough" to delete or correct the DS
Date Format in header was not correct
DVW Test: The method: With the sample in a well-heated condition (At the conclusion of the Normal Temperature Test). But the sample ID tested (1047A1-1) in DVW test was not same as sample ID tested (1047A1-2)in Temperature Test. Also the test date is not same (DVW in 10/17, but temperature test in 10/29). Not found the engineering judgment in DS

## Slide 4
Observation – from Cammy Hung
Project F857387 (VW-1 Test) – Testing staff didn’t record the distance between the upper surface of the cotton and the tip of the blue inner cone of the flame.
Requirements :
Clause 9.4F of SOP 00-LC-S0258, The product testing staff shall record test method parameters.

## Slide 5: Observation – from Shannon Tsui
Equipment:
Referenced Requirement - LC-P0031– Issue 13.0
    Verification data for validated Support Equipment listed in LEM shall be periodically confirmed through a designated maintenance schedule. (Cl. 8.4.4)
    In Service Verification shall be conducted according to a defined procedure. (Cl. 8.6.3)
Non-Conformance
    Verification and Maintenance Check Form for Equipment: 1 time per year for HWS005, 6 times per year for MF002
Objective Evidence
    Hardwood Surface HWS005 (Initial verification only)
    Micrometer MF002 (1 verification only)

## Slide 6: Observation – from Shannon Tsui
Power Quality Measurement:
Referenced Requirement 00-LC-P0025 – Issue 9.0, Cl. 7.1.1
                00-OP-S0029, Issue 5.0
    Validation required at least annually for single phase circuits up to:
      120V 60Hz up to 20A;
      230V 50Hz up to16A;
      240V 60Hz up to 15A.
Non-Conformance
    Power quality measurement was not conducted at 230 V.
Objective Evidence
    Equipment No: PS020 1417 (Power Bar No. 2, Socket 1), PS012 1419 (Power Bar No. C1, Socket 1)

## Slide 7: Observation - from Brian Wong
UL Mark Certification Program Policy
Referenced Requirement – ISO17025 & 00-OP-S0067 (issue 12)
  ISO17025 Clause 4.4 – Any difference between the request or tender and the contract shall be resolved before any work commences.
  00-OP-S0067 Clause 150.6.5J – The following statement shall be provided for investigations involving tests: “Test results relate only to the items tested”.
Non-Conformance
  Completion date is extended without any notification / agreement to customer.
  Missing statement “Test results relate only to the items tested” in test record. (issue)
Objective Evidence
  Projects 10NK80070 (3009B)

## Slide 8: Issue 1 – from Paul Ip
NSEP Training : New Staff – Molly Woo
Requirement: 81-TC-S0025
NC:
Not Found a training record for “Data Recording” in LMS, but found the training sign-up sheet
Not Found training evaluation result.
<Resolved>

## Slide 9: Issue 2 – from Paul Ip
Project 10NK06606
NC: Final reviewer have some comment in the final review. But not found the feedback from the Reviewer in training.
Project 10CA21658
NC: Not found the customer reply for the item 1 issue Prelim Construction Letter.

## Slide 10: Issue 1 - from Brian Wong
UL Mark Certification Program Policy
Referenced Requirement – 00-OP-S0067 (issue 12) & 00-LC-S0258 (issue 16)
  00-OP-S0067 Clause 150.6.5D – Identify each standard and correlate each test to the corresponding standard.
  00-LC-S0258 Clause 9.2C – Initial date and complete explanation should be provided if test method changed after preliminary review
Non-Conformance
  Test name in Test Record cannot be correlated to test list in Datasheet.
  Test method in datasheet is revised without initial, date and explanation.
  Typo error on model name in test record.
Objective Evidence
  Projects 10CA08882 (3009A)

## Slide 11: Issue 2 - from Brian Wong
UL Mark Certification Program Policy
Referenced Requirement – 00-OP-W0038 (issue 5)
  00-OP-W0038 Clause 9.1.3B – Test equipment information shall be recorded in the Meter Use tool in LPM database.
Non-Conformance
  Equipment ID: DC044 is missing in LPM. (Fixed)
  Typo error on electrical rating in CRD.
Objective Evidence
  Projects 10CA14896 (3015A)

## Slide 12: Issue 1 – from Shannon Tsui
Power Quality Measurement:
Referenced Requirement - ISO/IEC 17025: 2005(E), Cl. 5.2.3
    The laboratory shall use personnel who are employed by, or under contract to, the laboratory. Where contracted and additional technical and key support personnel are used, the laboratory shall ensure that such personnel are supervised and competent and that they work in accordance with the laboratory’s management system.
Non-Conformance
    No Training records to show Rita Wong’s technical competence
Objective Evidence
    PS017 1424 (Power Bar No. 3, Socket 1)120 V, 50 Hz
Quick Fixed
    Training record was updated in LMS

## Slide 13: Issue 2 – from Shannon Tsui
Sample:
Referenced Requirement - 00-LO-P0026 – Issue 7.0, Cl. 6.1.9
    Designated sample storage areas are strictly for sample storage and are not to be used for other storage purposes.
Non-Conformance
    Unlogged FUS samples were located next to Test Bench No. 1
Quick Fixed
    New designated sample storage area next to Pel Kwok’s cubical is assigned

## Slide 14
Issue 1 – from Cammy Hung
Project F744765 (water immersion Test) –During test demonstration, testing staff use Unit Conversion (Global ID 56386), but haven’t input into meter use.
Requirements :
Clause 9.4L of SOP 00-LC-S0258, The product testing staff shall identify validated calculations in the list of test equipment used per 9.3.4 and 9.3.5 C (e.g., change of resistance temperature calculations, unit conversions.)

## Slide 15: Issue 1 – from Simy Li
Referenced Requirement : 11-LO-W0400 Cl 9.1
Once the SM confirm the sample holder will leave the company, he / she …. has to transfer the samples to another staff
Non-conformance :
Staff who left the company still owned the samples according to LIMS -> No change of ownership
Evidence :
Dickson Lo (all under Dickson Lo hasn’t been changed the ownership) - Done
Sample Tag # 657519 still belonged to Cherrie Ip
Sample Tag #1023818-PRO marked return to VSIC but owned by Ernest Cheung (for proficiency test use)

## Slide 16: Issue 1 – from Kelvin Tsoi
DAP regarding the PPP, DAP co-ordinator needs to be aware of the following:
How to authorize PPP?
What is the flow of PPP?
How to link up the PPP profile in the DAP website?
Who is the qualified signatory?
Remaks :
Due to the PPP is too new, we will revisit those areas in next audit.

## Slide 18: THANK YOU