## Slide 1
Internal Compliance Audit – Closing Meeting
UL International Italia S.r.l (Carugate)

## Slide 2: Agenda:
Audit Scope
Technical Audits in Compliance Audit Process
Objectives and Focus Areas
Executive Summary
Risks and Raised Concerns
Recommended Improvement Roadmap and Next Steps
Assessment Team:
Lead Auditor: Michelle Howell
Auditors: Danny Van Der Waal, Grainne Cheevers, Przemyslaw Zaleski.

## Slide 3: Audit Scope
UL INTERNATIONAL ITALIA SRL testing and calibration activities including:
• ISO/IEC 17025:2017 calibration laboratory activities under CL-163 accreditation by IAS performed locally at UL International Italia Carugate.
• ISO/IEC 17025:2017 testing laboratory activities under 1363 A and 1363 B accreditation by Accredia performed locally at UL International Italia Carugate.
• ISO/IEC 17025:2017 testing laboratory activities under TL-553 accreditation by IAS performed locally at UL International Italia Carugate.
Commercial testing – HVAC, Lighting, Lab & Ops, Power & Automation, Small Appliances, CTECH – Safety, Appliances.
• IECEE CB Scheme laboratory activities under recognition no TL-155 by IECEE including: product evaluation; product testing.
• ISO/IEC 17065:2012 Conformity assessment — Requirements for bodies certifying products, processes and services (applicable requirements only)

## Slide 4: Audit Sampling
Areas sampled and assessed against relevant requirements:
  Organization and Leadership;
  Competency management,
  Handling of samples,
  Reporting of results,
  Evaluation procedures,
  Records management,
  Customer feedbacjk and Complaints,
  Non conforming work
  Management System Documentation.

## Slide 5
Confirming activities and scope reflects on ATLAS records
Merging ICA and technical audit to cover CAB regional schedule of accreditation within accreditation cycle
Assessing implemented management system on all stages of project delivery including evaluation, testing and reporting.
Verifying conformance with Standards, Accreditation and Scheme as well as UL requirements.
Focus on expectations, issues and objectives identified by local, regional and global Leadership Team.
Provide feedback to Partners to reassure the capability, efficiency and consistency of implemented business processes.
New ICA Report - Europe
Additional ICA deliverables in Europe are:
    Maturity level,
    One UL Integration level
    Governance level
    Associated Potential Risk linked the findings,
    Voluntary Suspension of Operations
    Impartiality
    Share Knowledge and best practices for improvements.
Executive Summary
Strengths & OFIs
Findings and Associated Risks
Governance, Management System Maturity Level and One UL Integration Indicators
Audit Result
Annual Improvement Roadmap
Objectives and Focus Areas

## Slide 6
Utilization
Technical Audits per service
Conformity Process implementation per COU
Complete Business Overview
ISO / IEC 17065:2012
Additional Business Requirements
ISO / IEC 17025:2017
Accreditation Requirements
Regional Requirements
UL Safety ITA
Local Test Witness
Remote Test
Witness (SME)
Local Project Review
Remote Project
Review (SME)
Completed by Regional ICA Team
Organised and / or completed by AMS-R
Completed by Regional IQA Team with input from COU
One UL Integration
Maturity Level
TAT
H&S incidents
Single Source of Data per CAB / COU
Throughout the year
Technical
Interview (SME)
Competency Monitoring / PTP

## Slide 7
CAB Capabilities:
https://ul.sharepoint.com/sites/GRPEMEALAEMEALAQualityTeam/Lists/CAB%20Capabilities/AllItems.aspx
CAB Standards / Schemes:
https://ul.sharepoint.com/sites/GRPEMEALAEMEALAQualityTeam/Lists/CAB%20Activities/AllItems.aspx
Storage and accessibility

## Slide 8: Risks and Raised Concerns
ISO 17025:2017 | Description of Finding | Expected outcome | Risk Level | CAR #
6.2.1 | Test was executed in Normal fault conditions, however the voltage limits were applied as for Single Fault Conditions (did not cover tolerances.) | Investigate to obtain root cause and implement corrective and preventive actions. | Medium | 01
7.11.2 | Validation records for software Integral 3.2 from Labsphere were not available. | Locate / obtain validation records for Labsphere. | Low | 02

## Slide 9
,
Laboratory Process Improvement and Maturity
Time
Improvement
Standard
Continuous Improvement
Consolidation through Standardization
Improve
Plan
Check
Do
Standard
I
D
C
P
5-10 Year
Vision
Step 1
Annual
Review
3-5 Year
Plan
Annual Plan
Development of
Initiatives (plan)
Progress
Control
Implementation
Step 2
Step 3
Step 4
Step 5
Step 6
Step 6
Improve
Plan
Check
Do
Laboratory services
Processes
OBJECTIVE
P1
D1
I2
C1
I1
C2
D2
P2
In
Cn
Dn
Pn
Process Maturity
Y1
Y2
Y3
Y4
Lab
Technology
P
D
C
I
P
D
C
I
CSE
P
D
C
I
CSA
P
D
C
I

## Slide 10
UL INTERNATIONAL Italia S.r.l – Laboratory Management System Maturity Level
Chart: EMC-W Lab: 2018
Chart: EMC-W Lab: 2020
10
8
6
4
Basic QMS IMPLEMENTATION
Accreditation & Scheme COMPLIANCE
BEST KNOWN METHOD
(BKM)/WORLD CLASS
CUSTOMER SATISFACTION
IMPROVEMENT
0
The assessment result of performance of implemented Quality Management System for UL International Italia S.r.l shows a good level of compliances with the standards and processes.
There are no high-risk areas that need urgent attention, 1 minor observation, 1 finding and many opportunities for improvement identified to further strengthen the processes and systems.
Chart: 
| General | Structural | Resources | Process | Management system
 | 8 | 8 | 7 | 7.8182 | 8

## Slide 11: Summary
Based on review of testing and evaluation methods, and interviews with the staff including test witnessing for CBTL, the assessment team concluded that the quality system of UL International Italia S.r.l meets the requirements of ISO / IECE 17025:2017 and 17065:2012 where applicable. Findings, observations and opportunities for improvement have been raised to further strengthen the system and ensure continued compliance.
Strengths
Laboratory management are positive and proactive in terms of improving and implementing the quality management system to achieve a high level of compliance. The attitude to the audit was very positive with laboratory management welcoming all opportunities for improvement.
Overall knowledge of the QMS, procedures and processes is very good. Laboratory management navigated the systems well and responsibilities were shared equally.
Knowledge of the accreditation process is excellent and as a result the accreditations are very well managed by the laboratory.
Weaknesses
Non conforming work.
Competence record storage
Performance of equipment and control of spreadsheets in the calibration lab.

## Slide 12
Next Steps
Integration to the enterprise manuals.
Continue to ensure ATLAS is up to date.
Review CAB Capabilities.
Implement internal audit plan.
Plan 2023 IQA audit.

## Slide 14: Attendees
Opening Meeting Attendees – 17th October
Colonna, Gianluigi
Howell, Michelle
Pozzi, Fabio
Zamarioli, Luca
Personeni, Davide
Tibaldini, Luca
Rivoltella, Dario
Scalisi, Manuela
Pedroni, Fausto
Atzeni, Davide
Fumagalli, Corrado
Ciori, Paolo
Pellegrino, Luigi
Van Der Waal, Danny
Alzate, GianCarloBattista
Redaelli, Matteo
Zaleski, Przemyslaw
Closing Meeting Attendees – 5th December
Zanichelli, Ursula
Coldani, Daniele
Maresca, Davide
Parmiani, Walter
Casati, Francesco
Howell, Michelle
Scalisi, Manuela
Rebecchi, Ermanno
Marengoni, Chiara
Bagheri, Iman
Tibaldini, Luca
Redaelli, Matteo
Rivoltella, Dario
Pozzi, Fabio
Alzate, GianCarloBattista
Fumagalli, Corrado