## Slide 1
Internal Compliance Audit – Closing Meeting -
2023-377-IQA, Calibration laboratory Carugate (ISO/IEC 17025)
IAS scope CL-163
Operationality QMS and Calibration activities

## Slide 2: Agenda:
Audit Scope
Non-Conformities
Opportunities for Improvement
Risks and Raised Concerns
Summary
Next Steps
Assessment Team:
Lead Auditor: Danny van der Waal

## Slide 3: Attendees
Opening Meeting Attendees – 28 Sept. 2023
Closing Meeting Attendees – 9 October 2023
Present:
Available:
Present:
GianCarloBattista Alzate, Laboratory Leader – Calibration;
Dario Rivoltella, LAB manager Carugate;
Manuela Scalisi, Quality Engineer Carugate.
GianCarloBattista Alzate, Laboratory Leader – Calibration;
Dario Rivoltella, LAB manager Carugate;
Manuela Scalisi, Quality Engineer Carugate;
Neil Friggi, Quality Eng. Manager EU

## Slide 4: Audit Scope
Reviewing implementation of the management system (compliance ISO/IEC 17025: 2017, Global Integration ULID-014435; integration of ULID-018470 and ULID-009203. Use of ATLAS)
Control calibration activities as recognized by IAS focused on Torque and Pressure Calibrations
Availability and use of PDCA mechanism
Calibration process performance by Torque demonstration via camera; vertical audits on performance of Pressure calibration/ order intake/ planning/ performance/reporting.
Integration of IAS regulation AC204, Evaluation of uncertainty, off site measurements, measurement traceability
Areas sampled and assessed against relevant requirements:
  Organization and Leadership in the lab; Competency; control of equipment; Risk management; Reporting results; Records management, maintenance and use of management system (par 8 ISO/IEC 17025) etc…...

## Slide 5: Non-conformities
CAR 1: For recording Monitoring data of personnel competences, it is required in procedure ULID 005934 to use Form-ULID-000611 (DCS:00-LO-F0861). To cover the requirements for monitoring on basis of on-going/ annual reviews, it is required to fill in the dates in the column “review date” per element registered in Form-ULID-000611. The form used for monitoring Antonio Iovino competence was demonstrated. In this form no (review) date of competences was noted. It is not possible to determine that the monitoring took place in an on-going/ annual process. Registrations were insufficient.

## Slide 6: Opportunities for Improvement
  OFI 1 :IAS policy for offsite calibration is required. Offsite measurements has to be formally accredited by IAS. This is not yet traceable on the scope for accredition CL-163.
  Offsite calibration requires at least checks on:
  Registration and storage of environmental conditions at the site;
  Proof that the measurement set up was checked on acceptability before starting the calibration (verification check on determined criteria);
  Determination of influences as defined in the CMC uncertainty budget;
  Offsite calibration is a method to perfrom certain calibrations outsite the laboratory. This mehod is not defined in a workinstruction;
  The current checklist for offsite calibration, does not prescibe these essential topics but should have if you want to be accredited for it.

## Slide 7: Opportunities for Improvement
  OFI 2 :
  The schedule for ILC’s/PTP shows a GAP for Weight, Torque, Force, Fluid and LUX.
  It is explained as impossible to cover all quantities with relevant PTP’s. It is recommended to organize intra lab comparisons. This means to select an object on the quantity to calibrate and get the reference values from the National Measuring Institute. After that, the object can be calibrated by different Calibration engineers (anonymously) and results coming from different persons can be compared (use the En formular) and judged on consistency. Next to a quality check, this can be interpreted as a competence check as well. The outcome gives an acceptable quality check on performance indicators of the specific calibration.

## Slide 8: Risks and Raised Concerns
ISO 17025:2017 | Description of Finding | Expected outcome | Risk Level | CAR #
OFI 1 | OFI1 for offsite calibration indicates a low level of assurance to the process. Expect the risk level for operational activities low due to the high level of calibration competency of the calibration engineers and its management. | . | Low | 
OFI 2 | OFI 2 initiative to create alternative quality checks if there are no ILC/PTP available is can not been demonstrated. However, when the lab participates, the score appeared to be excellent, so it is expected that the risk is low for quantities not reviewed with an ILC/PTP. |  | Low | 
CAR 1 | Believe for what is told that there is on-going interaction with personnel about competence monitoring, but there is not enough evidence to demonstrate trace back with current registrations. Medium risk. |  | Medium | 

## Slide 9: Summary
Based on review of testing and evaluation methods, and interviews with the staff, the assessment team concluded that the quality system of the Lab Warsaw meets the requirements of ISO / IEC 17025:2017 where applicable. Opportunities for improvement have been raised to further strengthen the system and ensure continued compliance.
    Strengths
    Experienced Calibration management (Filippo and GianCarlo)
    Very well documented system (locally and global)
    Excellent support on quality issues from higher lab management and AMS representative
    Visibility/ definition of IAS objectives / KPI / opportunities/ risks and structure to control them on effective implementation
    Weaknesses
    Process assurance for offsite calibration
    1 Calibration engineer left, and the number of calibrations increased with 20% ….. Where does this stop.

## Slide 10: Next Steps
Plan 2023 Audit
  Expectations and Dates:
  Start, plan and implement improvements on written OFI’s and CAR with your own PDCA cycle for OFI’s and CAR in the GCAR system (expect the quarterly management review sessions).