## Slide 1: Data Acceptance Program Detailed Requirements
Presented by:
XXXXXX
UL Knowledge Services, Division of UL LLC
Month XX, XXXX

## Slide 2: Instructor Introduction
Name
Department
Location
UL Experience
Education and other Work Experience
Notes: Instructor Notes:
Introduce your co-instructor, yourself, and others that are assisting or observing.
Introduce yourself and provide details as shown on the slide.
Add other pertinent background information that will help establish your credibility.

## Slide 3: Participant Introductions
Please tell us your name
Company and Location
Role
What are you expecting from the training?
Notes: Instructor Notes:
Have participants fill out their name tents and place where instructor can see them.
Have participants stand and introduce themselves as indicated on slide.
On flip-chart, write down what each participant wants to get out of training.
Post flip-charts.
Explain that we will come back to these expectations at the end of the training to see if they were met.

## Slide 4: Administrative
Emergency Procedures
Breaks and Lunch
Restrooms
Smoking
Pagers and cell phones
Notes: Instructor Notes:
Identify Emergency procedures as appropriate for the location.
Inform learners they will be taking a couple of breaks throughout the class.
Identify bathroom, cafeteria, and coffee station locations (include the location's rules for food consumption during training.)
Tell learners that they can leave the room when necessary.
Describe company smoking rules as appropriate.
Ask learners to turn off or mute their cell phones.
Remind learners that participation during the class is welcome and encouraged.

## Slide 5: Course Purpose
This course provides:
  Guidance for becoming a DAP Participant
  Implementation Details and Next Steps
  Detailed review of DAP requirements using ISO/IEC 17025 and associated DAP Client Guidance Documents
Notes: Instructor Notes:
Review purpose of the training.

## Slide 6: Course Map
Instructor Led Modules
eLearning Module Topics
What is DAP?
  The DAP Programs
  Authorized Signatories
  DAP General Requirements
  UL Knowledge Solutions
  eLearning Review Q&A
  DAP Requirements using ISO/IEC 17025
  Detailed Requirements Introduction
Notes: Instructor Notes:
Show how the module fits in with the full course.

## Slide 7: Questions
Notes: Instructor Notes:
Solicit and answer any questions learners may have about what we will be doing today.

## Slide 8: Course Map
Instructor Led Modules
eLearning Module Topics
What is DAP?
  The DAP Programs
  Authorized Signatories
  DAP General Requirements
  UL Knowledge Solutions
  eLearning Review Q&A
  DAP Requirements using ISO/IEC 17025
  Detailed Requirements Introduction
Notes: Instructor Notes:
Show how the module fits in with the full course.

## Slide 9: DAP Overview eLearning Review
Notes: Instructor Notes:
Introduce Module 2 Topic.
Explain that this module is a debrief for the eLearning that participants completed before the instructor led portion of the course, and that this is an ideal time to ask any questions about the topics contained within that module.

## Slide 10: Topic 1: What is DAP?
Notes: Instructor Notes:
Ask if anyone had any questions about this topic after completing the eLearning?

## Slide 11: Review
What are the main benefits of the DAP?
What is required by DAP participants to continue submitting test data generated at their lab, to UL?
Notes: Instructor Notes:
Solicit answers and discussion from the class for the questions shown on the slide.
1. Increased control over timing and schedule for product testing and certification
  Decreased turnaround time to complete the certification process
  More flexibility for clients.
2. Successful annual Assessments
BE SURE TO EXPLAIN: DAP participation is not accreditation to ISO/IEC 17025 and should not be considered as one. Rather it is a 2nd party audit developed by UL to allow for the acceptance of externally generated test data to support UL Mark certification decisions.
A successful audit means the DAP participant is compliant to “DAP requirements”, not 17025. Our audit does not cover all requirements in 17025.

## Slide 12: Topic 2: Types of Data Acceptance Programs
Notes: Instructor Notes:
Ask if anyone had any questions about this topic after completing the eLearning?

## Slide 13: Review
What program is typically a good starting point for clients wishing to become a DAP participant?
What are the five main DAP Programs?
Which two programs are typical entry programs for DAP participants?
Notes: Instructor Notes:
Solicit answers and discussion from the class for the questions shown on the slide.
1. Witness Test Data Program (WTDP)
2. Witness Test Data Program (WTDP), Client Test Data Program (CTDP), Total Certification Program (TCP), Third Party Test Data Program (TPTDP), Preferred Partner Program (PPP).
3. Client Test Data Program (CTDP), Third Party Test Data Program (TPTDP). Not Witness Test Data Program (WTDP): While DAP controls WTDP requirements, it does not monitor or audit clients using the program.

## Slide 14: Topic 3: Data Acceptance Program Requirements for Participation
Notes: Instructor Notes:
Ask if anyone had any questions about this topic after completing the eLearning?

## Slide 15: Review
Which three DAP programs require most ISO/IEC 17025 elements to be addressed?
In addition to ISO/IEC 17025 elements, what are some of the additional UL DAP requirements?
How will the annual assessments change after the Initial audit and you become a DAP Participant?
Notes: Instructor Notes:
Solicit answers and discussion from the class for the questions shown on the slide.
1. Total Certification Program (TCP), Third Party Test Data Program (TPTDP), Preferred Partner Program (PPP).
2. DAP Client Guides.
Authorized Signatories responsible for meeting DAP requirements and review of submitted data.
L56 Agreement
3. No longer full assessment. Sampling of tests for the Technical Audit and quality clauses assessed annually.

## Slide 16: Topic 4: DAP Authorized Signatory Responsibilities
Notes: Instructor Notes:
Ask if anyone had any questions about this topic after completing the eLearning?

## Slide 17: Review
What are DAP Authorized Signatories responsible for?
Notes: Instructor Notes:
Solicit answers and discussion from the class for the questions shown on the slide.
1. Attest to the assignment of qualified personnel to conduct testing
  Review, sign, date, submit completed data sheet packages to UL
  Understand and address NCRs with acceptable corrective actions plans
  Implement corrective action plans
  Notify UL of
    Signatory additions or deletions
    Changes in test lab location
  Request additions, deletions, or revisions to the DAP Scope of Participation.

## Slide 18: Topic 5: UL Knowledge Solutions for DAP clients
Notes: Instructor Notes:
Ask if anyone had any questions about this topic after completing the eLearning?

## Slide 19: Review
What are the three areas that UL offers assistance with for new and existing DAP participants?
Notes: Instructor Notes:
Solicit answers and discussion from the class for the questions shown on the slide.
1. Lab Support Services, Knowledge Transfer, GAP Assessment.

## Slide 20: Discussion and Questions
Notes: Instructor Notes:
Address any questions before moving on to the next module.

## Slide 21: DAP Requirements using ISO/IEC 17025

## Slide 22: Course Map
Instructor Led Modules
eLearning Module Topics
What is DAP?
  The DAP Programs
  Authorized Signatories
  DAP General Requirements
  UL Knoweldge Solutions
  Detailed Requirements Introduction
  eLearning Review Q&A
  DAP Requirements using ISO/IEC 17025
Notes: Instructor Notes:
Show how the module fits in with the full course.

## Slide 23: Module Objectives
Upon completion of this module you will be able to:
  Describe various ISO auditing terms
  State the purpose and scope of ISO/IEC 17025
  Identify components of the Management and Technical requirements in ISO/IEC 17025 used for DAP
Notes: Instructor Notes:
Briefly review the objectives.

## Slide 24: What is ISO/IEC 17025?
An international standard to judge competency of a testing and/or calibration laboratory based on:
  Personnel
  Facilities (physical location)
  Equipment
  Management system
Technical compliance
Quality management system in place
Ultimate goal:
Reliable Test Data
+
=
Notes: Instructor Notes:
Review the slide.
Indicate that ISO/IEC 17025 is a competency standard. When accreditation bodies conduct assessments, they are looking specifically for competency in the following areas of the client lab (read the list on the slide).
Discuss that a standard is a published document which sets out specifications and procedures designed to ensure that a materials, product, method or service is fit for its purpose and consistently performs in the way it was intended.
Also point out that standards define quality and establish safety criteria. Standards and conformance are keys to ensuring quality and consistency of various measurements.

## Slide 25: ISO Terminology
“Shall”
– Mandatory requirement (must be audited)
“Should”
– A suggestion, but not mandatory
“Notes”
– Demonstrates implementation of a process.
Example: a blank form is a document; a completed form becomes a record.
“Record”
– Provides clarification of text, guidance and examples
– Not audited
– Required physical written document
“Procedure”
Notes: Instructor Notes:
Review the slide. (Each definition will be shown individually upon click.)
Prompt learners to fill in their interpretation of the definitions in their learner guide.
A record is simply ‘who’ did ‘what’, ‘when’.

## Slide 26: Scope
Section 1: Scope
Testing Laboratory
  Evaluates/analyzes test samples
Calibration Laboratory
  Step 1 - establishes a relation between the quantity values with measurement uncertainties provided by measurement standards and corresponding indications with associated measurement uncertainties (of the calibrated instrument or secondary standard)
  Step 2 - uses this information to establish a relation for obtaining a measurement result from an indication
Some requirements pertain only to one type of lab.
Some labs do only one service and some do both.
Notes: Instructor Notes:
Review the slide.
The standard applies, as we have seen, to all kinds of labs, regardless of size. Many are commercial labs; some may be 1 or 2 persons working in a lab which is part of a manufacturing location.
Indicate that we have seen that the standard applies to all types of laboratories. However, there is recognition that some client labs do perform calibration on their own equipment.

## Slide 27: References and Terms and Definitions
Section 2: References
  ISO 17000 and VIM—vocabulary in metrology
Section 3: Terms and Definitions
  ISO 17000 and VIM—vocabulary in metrology
Notes: Instructor Notes: Review quickly and move on.
Use the following points for discussion:
  Terminology and other references are found in the International Vocabulary of Basic and General Terms in Metrology, or VIM.
  The ISO Standard is divided into 5 sections, with section 4 being the quality management requirements based on ISO 9001:2000 and section 5 being the technical requirements which are based on the old Guide 25 etc.

## Slide 28: Section 4: Management Requirements
4.1 Organization
4.2 Management system
4.3 Document control*
4.6 Purchasing services and supplies*
4.9 Control of nonconforming work*
4.11 Corrective action*
4.12 Preventive action
4.13 Control of records*
4.15 Management reviews
*Applicable For CTDP Participants
Notes: Instructor Notes: You’ll notice not all of the clauses in ISO/IEC 17025 are required by UL DAP.
Remind the class that these are the sub-sections or clauses that support the quality management requirements for 17025.
This section, provides the areas a lab should be competent in for carrying out testing and calibrations.
Emphasize that this section focuses on requirements for making sound management judgments.
As a result of complying with these requirements, the lab would be operating an effective quality management system.
Point out that we will review each clause in more detail in this module.

## Slide 29: 4.1 Organization / 4.2 Management System
4.1 Organization
Controls laboratory operation.
Ensures key lab personnel responsibilities clearly defined.
4.2 Management system
Quality documents drive consistency and continuous improvement.
Management’s commitment to good practice and quality of testing.
Notes: Instructor Notes:
4.1 Explain the purpose is to drive consistency of operation and control over the laboratory. Ensures that be defining lab roles, potential conflicts of interest are identified.
4.2 Explain management and quality documents drives consistency in the QMS.
(Optional)
Q: Why would 17025 cover organization and management? (Learners can fill this in in their guide.)
A: Because the DAP client lab is legally responsible for test data submitted to UL.

## Slide 30: 4.3 Document Control*
Procedures to control all internal and external QMS documents.
New and/or revised documents reviewed and approved
Master List of QMS documents maintained.
Periodic review of all QMS documents.
  Records of review and approval required
Procedures to describe changes to electronic documents.
Hand amendment of documents is allowed, but must be defined.
* Applicable for CTDP Participants
Notes: Instructor Notes:
Explain “QMS” – quality management system.
Master List typically prevents the use of invalid or obselete documents.
Whether a document is revised or not, there needs to be a record of periodic review.
REVIEW embedded example of doc control procedure to show something tangible (reset the audience’s clock!)

## Slide 31: 4.6 Purchasing services & supplies*
Records required to:
  Inspect services (calibrations) and consumables to meet requirements before use in lab.
* Applicable for CTDP Participants
Notes: Instructor Notes:
Review the slide.

## Slide 32: 4.6 Purchasing services & supplies*
Slide 32
Records required to:
  Inspect services (calibrations) and consumables to meet requirements before use in lab.
Purchasing Procedure
Purchasing Record
Notes: Instructor Notes:
Open and show the example documents.
Note that these are examples of what a purchasing procedure may look like and also a purchasing record.
Q: Ask participants if they have other ideas on what should be looked at to show compliance to 4.6
A: Actual inspection records for consumables. Calibration certificates to show equipment is in-tolerance and meets requirements.

## Slide 33: What to Inspect for Consumables
May include:
External Certificate of Conformance (C of C)
Identification (catalog #, model #, etc.)
Description of materials and detailed test method specifications
    For chemicals, should include both brand and generic names
Expiration date or shelf-life
Special storage requirements
00-OP-C0033, 00-OP-C0037
Notes: Instructor Notes:
Facilitate discussion on examples of test method specifications so they understand what is required on a consumable verification record.
If you have internet access, go to the public website to show clients the technical guide documents for the first time: http://www.ul.com/global/eng/pages/offerings/services/programs/dap/tools/
If no internet access, be sure to download the most current versions of 00-OP-C0033 (Critical Consumables) and 00-OP-C0037 (TCs) to show the class.
To get a better understanding of TC client guide, 00-OP-C0037, conduct class Activity 1 on next slide (and Activity 2 if time allows):
IF class is still unclear about TCs, review “TC Validation” from the zip file.

## Slide 34: Activity 1: Purchasing Services and Supplies
Directions:
  Review situation and sample materials in the DAP Activities handout.
  Refer to Clause 4.6 of ISO/IEC 17025 and 00-OP-C0037.
  Can your thermocouple junctions be used for test? Why or why not?
Time:
  Individual – 5 minutes
  Class debrief – 10 minutes
Notes: Instructor Notes:
Review the Instructor Notes
Review the directions for Activity 1 (below). We are reviewing TC wire. If the content is not clear to participants, and time allows, conduct Activity 2.
Tell the participants that they have 5 minutes to complete this task before discussion.
Refer participants to the DAP Purchasing Services and Supplies activity in their DAP Activities Handout for the “test protocol” and the purchasing document. If time allows, conduct Q&A with class.
Optional TC Activity 2:
Read the following scenario:
The lab supervisor is responsible to order his own lab consumables. He’s was running low on TC wire so he placed an order to one of the local electrical distributors (AME Electrical Supplies) for the TC wire. The quantity he ordered was one spool of TC wire which has 1000-ft length. The next day the spool of TC wire was delivered. Upon receipt, he immediately put it into service and asked one of his technicians to start making junctions for inventories. The tipping was done by twisting the pair of TC wire together. The lab’s practice is to reuse as many TC joint assemblies as possible without re-tipping and the newly made joint assemblies were mixed in with the used ones. The lab tech can just pull any of these junctions when needed to do UL testing.
Discuss these Questions & answers (given below each question) as a class:
  1. Do you have any concerns about this scenario?
    a. Twisting is not recommended.
    b. Mixing new & used assemblies
    c. Also see next questions
  2. What questions come to mind?
    a. How was the spool identified?
    b. Is the vendor approved?
    c. Did the wire come with a C of C?
    d. Was the wire & the C of C inspected before use - is there a record?
    e. Was the junction verified – is there a record?
A key teaching point is that two verification records are required: one for verification before use (like any lab consumable) and verification of TC junctions. (Whether assembled internally or externally).
Review the answer:
  Two verification records are required: one for verification before use (like any lab consumable) and verification of TC assemblies (whether assembled internally or externally).
  Validation result for 104 degree range, at 105.3 is out of spec. Suggest that it might be good to add a “comply/not comply” box to your form to help highlight results and take corrective action as needed.
  A key teaching point is that the process to confirm one critical consumable can be time-consuming.

## Slide 35: What to do if the Consumable Does Not Meet Requirements?
Contact supplier to resolve issue and replace consumable
Determine if consumable can be provided by UL
Look to obtain consumable from UL suppliers
Contact UL and obtain approval first if there is a need to perform the test using the current consumable
KNOWLEDGE CHECK:
What is required for critical consumables?
Notes: Instructor Notes:
Review the slide.
The client lab is responsible to obtain and use the proper critical consumable test materials.
Client should discuss with UL if consumables don’t meet requirements before performing the test and submitting data to UL – e.g. what the client should do.
KNOWLEDGE CHECK: What is required for critical consumable?
  Lab should have the test standards to show specifications
  Receipt and/or Verification records that show specifications are met
  Storage of consumables

## Slide 36: 4.9 Control of nonconforming testing and/or calibration work*
Procedure when test and/or calibration work does not conform to lab’s own procedures or customer requirements.
Notify UL when nonconformity could effect the validity of data submitted!
Corrective Action (4.11) is followed when nonconformities can recur.
ASK: If the product fails a test, is that nonconforming work?
ASK: What is potential nonconforming test work?
* Applicable for CTDP Participants
Notes: Instructor Notes:
Discuss the difference between non-compliance to a test vs. non-conformance to DAP or the lab’s own requirements.
ASK: What is potential nonconforming work?
EXAMPLES:
    Incorrect/incomplete recording of data.
    Equipment out of calibration/tolerance
    Equipment found to give suspect data
    Equipment malfunction
    Consumables used after shelf-life or expiration date
    Consumable re-verification is found out of tolerance
    Environmental conditions out of tolerance
    Observe inconsistent methods used during test
    Solicit more answers to get the class to understand.

## Slide 37: 4.11 Corrective Action*
Procedure to address:
  Nonconforming work
  Departures from policies & procedures
Start with root cause investigation.
Implement appropriate corrective actions.
Evaluate and monitor effectiveness of corrective actions taken.
Records of all actions including effectiveness check.
* Applicable for CTDP Participants
Notes: Instructor Notes:
Review the slide.

## Slide 38: 4.12 Preventive Actions
Pro-active process to identify opportunities for improvement.
Procedure to address:
  Initiation of preventive actions
  Apply controls to ensure effectiveness
Action plans developed, implemented and monitored.
Notes: Instructor Notes:
Review the slide.

## Slide 39: 4.13 Control of Records*
Procedure to manage, protect, back-up quality and technical records.
Original observations must be retained.
Retention times identified for all records.
    UL requires 5 years from date of record.
Test records contain enough information to repeat test as original.
    00-OP-C0025
  * Applicable for CTDP Participants
Notes: Instructor Notes:
Review the slide.

## Slide 40: 4.15 Management Review
Procedure to periodically review laboratory’s:
  management system
  test/calibration activities
Resulting actions carried out per agreed upon timeline.
Records of items reviewed, findings issued and actions taken.
Notes: Instructor Notes:
Review the slide.

## Slide 41: Section 5: Technical Requirements
*Applicable For CTDP Participants
5.2 Personnel*
5.3 Environmental conditions*
5.4 Methods and method validation*
5.5 Equipment*
5.6 Measurement traceability*
5.8 Handling of test/calibration items*
5.10 Reporting the results*
Notes: Instructor Notes:
Review the slide.
Section 5 of the ISO 17025 Standards discusses Technical requirements.

## Slide 42: 5.2 Personnel*
Maintain records of:
  experience/skills
  training
  qualification (authorization date)
Competency of personnel to:
  Perform tests/calibrations
  Operate equipment (5.5.3)
  Supervise trainees
* Applicable for CTDP Participants
Notes: Instructor Notes:
Review the slide.

## Slide 43: 5.3 Accommodations & Environmental Conditions*
Appropriate facilities for testing
  Environmental conditions
  Energy sources
Separation when neighboring testing/activities is incompatible
Controlled lab access, as needed
00-OP-C0035, 00-OP-C0036
* Applicable for CTDP Participants
Notes: Instructor Notes:
Review the slide.
Review client guide documents with class (internet or current version saved): 00-OP-C0035 (Environmental Conditions) and 00-OP-C0036 (Power Quality).

## Slide 44: 5.3 Accommodations & Environmental Conditions*
Monitor, control & record
  Typically, there are meters in each lab to measure temperature, humidity, and barometric pressure.
  Stop activities when environmental conditions jeopardize results.
  UL approval is needed to accept data if conditions do not meet stated requirements during test.
Notes: Instructor Notes:
Review the slide.
Explain that they would need to check ambient conditions in their lab before beginning testing.
Ambient conditions are not limited to just temperature, humidity or barometric pressure. It could be other factors that could affect the test results such as EMC, noise, light, air movement, etc.
Remind staff that this is one of the conditions of the Accuracy Method.

## Slide 45: 5.4 Test & Calibration Methods & Method Validation*
Have access to all standards in DAP Scope of Participation.
  Method meets customer needs (UL)
  Published Standards
  Valid Edition
  Supplemented where necessary (work instructions)
Deviation from test method only if documented and UL approved.
Records available if the following apply:
  Data transfer checks
  Calculations
  Validation of Lab-developed Software
Procedures for data back-up
* Applicable for CTDP Participants

## Slide 46: 5.5 Equipment*
Laboratory has necessary equipment
    Equipment rented or borrowed meets DAP requirements.
Meets accuracy requirements
    If not specified in standard, meets 00-OP-C0034
Calibrated/checked before use for values affecting results (range)
Used by authorized personnel
Procedures for safe handling, transport, storage, use, maintenance
Records of identification, accuracy check, instruction manuals, calibrations, damage, repair.
00-OP-C0034
* Applicable for CTDP Participants
Notes: Instructor Notes:
Review the slide.
Open 00-OP-C0034 to review equipment accuracy.

## Slide 47: 5.5 Equipment *
Damage control
Calibration status on equipment
Check before use for tests if goes outside of lab control
Procedure for intermediate checks, as needed
  Applies when equipment is one-time calibrated.
Procedure for correction factors, as needed
Safeguarded from adjustments
    Hardware
    Software
* Applicable for CTDP Participants
Notes: Instructor Notes:
Review the slide.

## Slide 48: 5.6 Measurement Traceability*
All equipment calibrated before use
Procedure and schedule required
External calibration services are:
    17025 accredited and need to provide logo on calibration certificates, OR
    NOT 17025 accredited and are evaluated by the DAP participant to show competence, measurement capability, traceability to SI units.
Calibration records include measurements results and uncertainty.
  00-OP-C0032 (Calibration Certificate analysis)
  00-OP-C0038 (Non-accredited Calibration labs)
  00-OP-C0045 (Calibration Intervals)
* Applicable for CTDP Participants
Notes: Instructor Notes:
Review the slide.
Open 00-OP-C00## docs to review various calibration requirements.
EXPLAIN when a lab can use a non-accredited calibration supplier.

## Slide 49: 5.6 Measurement Traceability*
DAP Participants can choose to perform In-House Calibration of equipment used to test for Scope of Participation
00-OP-C0038 (In-House calibration labs)
  Applicable ISO/IEC 17025 clauses: 4.3, 4.6, 4.9, 4.11, 4.13, 5.2, 5.3, 5.4, 5.5, 5.6, 5.10
  DAP conducts “audit within an audit” for the calibration lab.
  What information is required to be on an
  In-house (internal) Calibration record?
* Applicable for CTDP Participants
Notes: ASK: What information is required to be on an In-house (internal) Calibration Certificate?
  Everything that is required on an external calibration certificate (00-OP-C0032) EXCEPT that records may be kept in several locations (i.e MU records are separate from calibration report); Does not all have to be on 1 report/certificate.
  No need for accreditation logo.

## Slide 50: Accreditation Endorsements – Example of some ILAC MRA signatories, 00-OP-C0032
Argentina | OAA | www.oaa.org.ar
Australia | NATA | www.nata.asn.au
Austria | BMWA | www.bmwa.gv.at
Belgium | BELAC | www.belac.fgov.be
Brazil | CGCRE/INMETRO | www.inmetro.gov.br
Canada | CAEL | www.caeal.ca
Canada | SCC | www.scc.ca
China, Hong Kong | HKAS | www.itc.gov.hk/hkas
China, Peoples Republic of China | CNAS | www.cnas.org.cn
Chinese, Taipei | TAF | www.cnla.org.tw
Croatia | HMD | www.hmd.hr
Cuba | ONARC | www.onarc.cubaindustria.cu
Denmark | DANAK | www.danak.org
Egypt | NLAB | www.nlab.org
Finland | FINAS | www.finas.fi
France | COFRAC | www.cofrac.fr
Germany | DACH | dach-gmbh.de
Notes: Instructor Notes:
Review the slide.
Refer the class to the next pages in the DAP Guide.
Review information about accreditation endorsements.

## Slide 51: Activity 2: Reviewing Calibration Certificates
Directions:
Review materials in the DAP Activities handout.
  A Calibration Certificate is reviewed by your lab staff for a voltmeter used to test a plasma generator.
Use 00-OP-C0032 to assist in your analysis.
In teams, determine whether it meets the technical requirements and answer these questions:
  Can the voltmeter be used to gather data? Why or why not?
  What did you find out about the calibration data?
  What did you find out about the calibration vendor’s accreditation?
In teams – 10 min
Debrief as a class – 10 min
Notes: Instructor Notes:
Review the slide.
Give the class 10 minutes to discuss/prepare answers to questions.
If Private workshop at client’s facility, ask to use one of their certificates for a real-life example of this activity.

## Slide 52: 5.8 Handing of Test & Calibration Items*
Slide 52
Procedure for transport, receipt, handling, protection, storage, disposal of test samples.
Choose identification of samples so:
    Items cannot be physically confused
    Unique enough to correlate across records
    Accommodate subdivision of groups of samples
    ID is retained through life of item in lab
Test sample receipt control
    Issues reported to customer (UL).
    Record abnormalities, resulting discussions and next steps.
* Applicable for CTDP Participants
Notes: Instructor Notes:
Review the slide.

## Slide 53: DAP Technical Records: Clauses 4.13.2, 5.10
Original observations
Sufficient records maintained for audit trail and repeatability of tests
Identify person(s) responsible for test and checking results
Recorded at time observation/data is made
Mistakes crossed out, initialed and dated; not deleted
Changes initialed, dated, and justification
Notes: Instructor Notes:
Review the slide.
The topic of original observations and recording original observations at the time they are made should be introduced here (next paragraph). Also, error corrections* should be introduced here.
Technical Records will be described in more detail in the Data Recording and Reporting section of Module 5.
*Clause 4.13.2.3 addressed error correction. UL trains our staff to use "single line, initial and date" (or electronic equivalent) for error correction.
For changes, justification needs to be recorded, as need. For example, test deviations need to be L3 approved by UL.

## Slide 54: DAP Technical Records - continued
Data Transfers and Calculations checked in a systematic manner1
  Transfer of data from notebooks or chart recordings to final test report
  Manual calculation
Lab-developed software validated for use2
  LabView
  Excel Calculations
1ISO/IEC 17025 requirement, clause 5.4.7.1
2ISO/IEC 17025 requirement, clause 5.4.7.2
Notes: Instructor Notes:
Review the slide.
Technical records for original observations to documented datasheets need to be created. This process is called data transfer.
When UL conducts DAP audits, we check the intermediate steps to transfer the data from original observation to the datasheet. For example, if the original observation was recorded on a checklist before being entered on the datasheet, you must show evidence of this process and review of the data transfer, preferably by someone different than the person performing the transfer.
We also check that any software or excel calculations created by the DAP client, not by OEM, has been validated to produce repeatable/reliable data. Evidence of validation is required (i.e compare manual calculation to excel calculation).

## Slide 55: DAP Reporting Results: Requirements
Data Sheets/Test Reports shall contain*
  Title
  Name and address of lab conducting tests
  Unique identification of test report, each page, and “end”
  Name and address of customer
  Identification of items tested
  Receipt date of test item where critical to validity of results
  Date of performance of tests
  Sampling plan, where applicable
  Results with units of measurements, where applicable
  Name, function and signature of persons authorizing test report
  Statement that results apply only to item tested, where relevant
  *Refer to ISO/IEC17025, clause 5.10.2 for full text of above and refer to
  00-OP-C0025 for complete list of items required by UL.
Notes: Instructor Notes:
Review the slide.
Remember that UL considers data sheets to be part of the test report.
Emphasize that the completed data sheet package is the sum of all the DAP participant’s processes and activities, in other words the “product”, or output of the process.
State the importance that the DAP participant needs to ensure that completed data sheet package meets the DAP Data Recording and Reporting requirements in 00-OP-C0025.

## Slide 56: DAP Reporting Results: Additional Reqs
Data Sheets/Test Reports shall contain*
  Testing staff signature
  Correlation to equipment
  Environmental conditions (whether or not there is a location on data sheet)
  Deviations in methods or conditions
  Calculations (including unit conversions)
  Other information needed by UL for data interpretation (photos, drawings, etc.)
  Opinions and interpretations
Test reports not complying with 00-OP-C0025 will be returned for rework.
  *Refer to ISO/IEC17025, clause 5.10.2 for full text of above and refer to
  00-OP-C0025 for complete list of items required by UL.
Notes: Instructor Notes:
Review the slide.
Re-state the importance that the DAP participant needs to ensure that completed data sheet package meets the DAP Data Recording and Reporting requirements in 00-OP-C0025. Otherwise, the completed data sheet package will be returned to the lab as non-conforming. Numerous key DR&R steps all add up to the final “product” (i.e. getting data sheets approved by UL , getting deviations approved by UL).

## Slide 57: Key Components of the Data Sheet Package
Cover pages
Test data sheet pages
00-OP-C0025
Notes: Instructor Notes:
Review the slide.
Stress to the participants that the data sheets on the next couple of slides are just examples and not to be inferred as templates.
Review the main components of the data sheet package with the class.
Also refer the class to the Data Recording, Reporting and Related Requirements, 00-OP-C0025.
IMPORTANT: Clarify the Process to obtain current UL data sheet packages

## Slide 58: Data Sheet Package: All Pages
  Applicant File Number (optional)
  Project Number (or equivalent unique identification)*
  Page Number – sequentially numbered pages (suggest page X of Y format).
  Date – use YYYY-MM-DD with exceptions.
  *Contact project handler for project number. If using own unique identifier, must correlate data sheets to a UL Project number, to ensure traceability.
Notes: Instructor Notes:
Review the slide.
IMPORTANT: Tell the class to refer to the Data Sheet Example in their DAP Activities to see the various sections more easily.
Also refer the class to the 00-OP-C0025.
Clarify the following:
  Every page in the data sheet package must contain the Applicant File Number, UL Project Number or unique ID (see below), and Page Number at the top of the page.
  The client is advised to get the project number in advance and put on datasheets. If not available, a unique ID of the data sheet package must be recorded and correlated to the UL project number.
  Be sure to tell them the UL date format is YYYY-MM-DD and point out the date format exceptions in 00-OP-C0025 “The issue date and the latest revision date of a test Standard or for test equipment calibration records provided the date format is not ambiguous”. Point out Date: 8/13/2007 is incorrect date format.

## Slide 59: Data Sheet Package: Cover Page
Total number of pages
Test location (Name and full address)
Name of applicant
Applicable test standard(s) by identification number, edition or issue date and revision date
List of tests
Special instructions (as applicable)
Notes: Instructor Notes:
Review the slide.
IMPORTANT: Tell the class to refer to the Data Sheet Example in their DAP Activities to see the various sections more easily.

## Slide 60: Data Sheet Package: List of Tests for Multiple Standards
If the product is investigated to more than one standard:
  Identify standards
  Correlate list of tests performed to each standard
1 | ✓ | TRANSFORMER OVERHEATING TEST (Mains supply transformers Overheating, CAN/CSA 22.2 No. 601.1): | Same test with different name.
Notes: Instructor Notes:
Review the slide.
If the product is investigated to more than one standard, such as Canadian, IEC or multiple UL standards, the standards shall be identified and the list of tests performed shall be correlated to each standard.
Review the slide and this section of the sample data sheet.
Add any information you feel is important.

## Slide 61: Data Sheet Package: List of Tests for Multiple Standards – continued
If titles of related tests among involved standards are different:
  Correlate between the titles
  Identify applicable clauses from each standard
TRANSFORMER OVERHEATING TEST: | Sec. 14
Mains supply transformers Overheating (CAN/CSA C22.2 No 601.1): | 57.9.1 a)& 57.9.1 b)
Notes: Instructor Notes:
Review the slide.
If the titles of related tests among the standards involved are different, correlation between the titles shall be provided and the applicable clauses from each standard shall be identified.
Review the slide and this section of the sample data sheet.
Add any information you feel is important.

## Slide 62: Data Sheet Package: List of Tests for Multiple Standards – continued
When a test method in one standard is representative of test method(s) in other standard(s):
    Record representation clearly under test method used.
    Correlate between single test performed and the test(s) it represents.
[x] The following tests conducted in accordance with UL 60601-1 were considered representative of the same tests required by Canadian Standard, CAN/CSA-C22.2 No. 601.1-M90.
Notes: Instructor Notes:
Review the slide.
When testing in accordance with the requirements in one standard is considered representative (exact or worst-case scenario) of the applicable requirements in another standard, this representation must be clearly recorded under test method.
Review the slide and this section of the sample data sheet.
Add any information you feel is important.

## Slide 63: Data Sheet Package: Additional Required Information
Environmental condition statements:
    “Unless specified otherwise in the individual Methods, the tests shall be conducted under the following environmental conditions. Confirmation of these conditions shall be recorded at the time the test is conducted.” (followed by the conditions).
    Or indicate, “No general environmental conditions are specified in the Standard(s)or have been identified that could affect the test results or measurements.”, if applicable.
Notes: Instructor Notes:
Review the slide and this section of the sample data sheet.
Add any information you feel is important.

## Slide 64: Data Sheet Package: Additional Required Information – continued
Sample information
  Unique ID, especially when multiple samples of same model/type are tested
  Sample source
  Clear correlation to test results
Notes: Instructor Notes:
Review the slide.
Review this section of the sample data sheet.
Add any information you feel is important.

## Slide 65: Data Sheet Package: Test Equipment
Identify equipment used for test:
    Calibrated instruments
    Non-calibrated test support resources (impact balls, probes, alcoves, etc.)
Unique identification
    Identify source of equipment if any is rented
Instrument type
Correlation to test
Range or function
Last and next calibration date
Notes: Instructor Notes:
Review the slide and this section of the sample data sheet.
As you refer to this information in the data sheet, step them through the following items:
  Instruments
  Test support resources – impact balls, probes etc.
  Instrument type
  Unique identification – if both UL and your own equipment are used, the equipment ID shall clearly identify the source of the equipment.
  Range or function
  Last and next calibration date
Point out “Verify Before Use”
  For equipment marked “Verify Before Use”, list any calibrated equipment used for verification purposes and record the verification data.

## Slide 66: Data Sheet Package: Test Equipment Example
Notes: Instructor Notes:
Review the slide and this section of the sample data sheet.
Add any information you feel is important.

## Slide 67: Test Methods
Deviations need to record:
    Technical reason
    UL Reviewer & date of approval
    Lab staff initials & date
  “The location of the transformer made this measurement very difficult to make. Jack Todd and Jim Smith of UL on 8/13/12 confirmed that the cold resistance of the winding measurement was not required by the Standard and did not need to be provided. SG 2012-08-13
Complete all test results sections in the data sheet
Unused fields marked “N/A”, sections lined-out, or “X” across page, initial and date
Notes: Instructor Notes:
Review the slide.

## Slide 68: Test Methods – continued
Record actual values.
When pre-printed results are provided, select applicable response.
NOTE: Strikethrough, in this case, does not require initials or date.
  Example:
    [X] There was no indication of breakdown.
    [ ] Breakdown occurred.
  OR,
    Breakdown [ occurred ] [ did not occur ].
Check limits and results obtained to determine if test passes or fails.
Notes:     Instructor Notes:
Review the slide.
Discuss Recording of ACTUAL Values.
Statements such as the following are NOT acceptable:
  The quantity was less than the maximum permissible value,
  More than the minimum permissible value, or
  A summary statement, i.e. no breakdown occurred to cover all the test instances.

## Slide 69: General Requirements for Recording Data
Record accurately, clearly and objectively at time data is obtained.
Indicate units for numerical data.
Label intermediate data with:
    test title
    date or a period of test
    range, speed
    file number
    project number
    page number
Sven Golly
Temperature (0C)
Time ( seconds)
Notes: Instructor Notes:
Review the slide.
Intermediate data: chart recordings or instrument printouts. Point out Test Date: 8/13/2007 is incorrect date format.

## Slide 70: General Requirements for Recording Data – continued
Record date of test on every page with data.
  Record start and end date, where applicable
Use black or dark blue ink if manual recording
If electronic,
Do not use highlighting, shading or background coloring
Minimize margin notes
Do not use embedded objects unless information contained in object is fully visible without opening it.
Notes: Instructor Notes:
Review the slide.
Color inks may NOT be used unless UL has verified the color will scan appropriately and has given you approval to use the color ink.
Embedded objects shall not be used in electronic datasheets since the information contained in the object is not included in the page count.
  Others reviewing the datasheets might not have the software necessary to open the object and if the datasheets file format is converted for storage on UL DMS, the conversion may identify the object as a graphic and the data is lost.
  Exception: Embedded objects are permitted if the information contained in the object is fully visible without having to open the object.
Add any other information you feel is relevant.

## Slide 71: Making Changes or Corrections
Neatly line out error
Record correct information
Initial and date the change
Provide reason for the change
Do not obliterate original text
  NO erase marks, correction tape, liquid,
  NO highlights, overwriting
SG 2007-08-13
850
Notes: Instructor Notes:
Review the slide.
Point out 8/13/07 is incorrect date format.

## Slide 72: Making Changes or Corrections – continued
Electronic datasheets
  Use “strike-through” format in MS Word
  Do NOT use “track changes”
Insert text directly or use “insert Text Box” to add explanations.
  Other text is not covered
  Do NOT use “Insert Comments” feature
Example:
Maximum Voltage 5 vdc
4.98 vdc
RAF 2005-10-26
Used more
accurate voltage
Notes: Instructor Notes:
Review the slide.
IMPORTANT: Tell the class to refer to the Data Sheet Example in their DAP Activities to see the various sections more easily.
If data is recorded electronically, error correction shall be the electronic equivalent to that in by use of “strikethrough” and addition of initials and reason for the change if the reason is not obvious, rather than deleting or replacing the erroneous data.
Using the track changes feature in Microsoft Word is NOT sufficient to meet this requirement. This feature may be used to facilitate review of the datasheet pretest but is not to be relied on for recording error corrections. Why? Because if changes are “accepted”, original data will be obliterated and ISO/IEC 17025 does not allow it.
Add any other information you feel is important.
For Example:
  RAF 2005-10-26 used more accurate range
  Maximum Voltage: 5 vdc
  Measured Voltage
  [strikethrough “5.1 vdc”] 4.98 vdc (RAF 2005-10-26 used more accurate range)
  Insert a text box and enter the information or type it adjacent to the erroneous data.

## Slide 73: Repeat Testing
Test repeated due to non-compliant results:
  Note non-compliance
  Sample information identifies revised sample construction for repeat test
  Reference repeat testing
Example:
“Noncompliant test results, see ___ for testing of revised construction.”
Tests repeated for any other reason:
  Use error correction method (single line, initial, date)
  Note reason
Notes: Instructor Notes:
Review the slide.

## Slide 74: Nonconforming Results
Nonconformance occurs when the test does not meet requirements of the test method in standard (different from non-compliance)
In the case of a nonconformance:
  Contact the UL Project Handler (PH) and Reviewer (L3)
  The Reviewer should review and approve acceptance of the deviation before testing proceeds
  Record changes for any deviation from the test method or program
  Initial, date, and provide the name of the PH or L3 who authorized the change
Examples:
  Ambient Conditions out of the range specified
  Power Outage
Notes: Instructor Notes:
Review the slide.
A nonconforming result is a failure to follow the test procedure, whereas a noncompliance is a failure of the product.

## Slide 75: Responsibilities of DAP Participant Staff
Ensure DAP datasheets meet document control (4.3). record control (4.13), test report (5.10.2) requirements, and 00-OP-C0025.
Record all test parameters on data sheet:
    Samples with correlation to test
    Test equipment with correlation to test
    Equipment information
    Environmental conditions
    Data recorded at time of observation
Conduct tests in accordance with standard.
Ensure intermediate data is properly labeled.
Proper error corrections.
Notes: Instructor Notes:
Review the slide.

## Slide 76: DAP Authorized Laboratory Signatory
DAP labs designate and document Authorized Signatory for COMPLETED data sheet packages sent to UL1
The Authorized Signatory’s signature and date on cover page attests that entire datasheet package meets 00-OP-C0025 and all other applicable DAP requirements2.
A client lab may have more than one Authorized Signatory.
00-OP-C0401
2ISO/IEC 17025 requirement, clauses 4.3, 4.13, 5.10.2
1ISO/IEC 17025 requirement, clause 5.2.5
Notes: Instructor Notes:
Review the slide.
The lab must define and document the personnel who will be the Authorized Signatory for COMPLETED data sheet packages being sent to UL.
The signature of the Authorized Signatory on the cover page of the data sheet package attests that the entire data sheet package meets all applicable UL, DR&R, ISO/IEC 17025, and DAP requirements.
We will discuss more in the next module and review client guide 00-OP-C0401.

## Slide 77: Activity 3: Data Recording and Reporting
Directions:
Divide into small teams. See DAP Activities handout.
You work at ABC Corporation CTDP lab.
You, the lab authorized signatory, are requested to review datasheets before submission to UL.
Use 00-OP-C0025 to determine the following:
Is any information on the data sheet package missing? What?
Is all data recorded correctly?
Team – 30 minutes
Class Debrief – 15 minutes
Notes: Instructor Notes:
Review the directions. Tell the participants that they have 30 minutes to complete the activity before group discussion.
Indicate that this activity focuses on reviewing the data sheet package in DAP Activities handout.
Emphasize that they are to determine whether any information in the data sheets is missing and if everything is recorded correctly.
Stress they should use the 00-OP-C0025 to evaluate the data sheets.
Ask if they have any questions.
For PRIVATE workshops, ask to review the lab’s datasheets for a more real-life look at DR&R requirements.
Monitor the activity.
Conduct a debrief of the activity. Be sure to ask different teams different questions.
Refer to your answer key, m4a1 answer key in the Instructor Guide for answers to this activity.
Ask:
  Is any information missing? What?
  Is all data recorded correctly? If not, what specifically is wrong?
  Did anyone find a section where comments needed to be added?

## Slide 78: Data Recording and Reporting Review
Key components of cover pages and data sheets
How to complete the data sheets correctly
Responsibilities in recording data during test
Making changes or corrections
Datasheets sent to UL that do not comply with DR&R requirements will be returned to the DAP Participant for correction.
Notes: Instructor Notes:
Review the slide.
Remind the participants to use the DAP Guides.
Re-emphasize the last bullet point - Nonconforming test packages sent to UL will be returned to the CTDP lab.
Add any additional comments you feel are important.

## Slide 79: Data Recording and Reporting Review
Key to DAP Data Recording and Reporting requirements:
  Test methods and validation
  Provide complete information
  UL authorized deviations
  Control of data, documents, records
Data recording and reporting errors is “non-conforming test work” as defined by ISO/IEC 17025, clause 4.9.
The Authorized Signatory is the final review for data sheet package submitted to UL
Notes: Instructor Notes:
Review the slide.
Ask participants if they have any questions regarding content covered in Module 5.

## Slide 80: DAP Requirements Summary
ISO/IEC 17025 addresses both testing and calibration lab services.
UL developed client guide documents to ensure consistent application of technical requirements.
UL’s adoption of ISO/IEC 17025 into the Data Acceptance Program is required to meet accreditation criteria.
Notes: Instructor Notes:
Review the slide.
As you can see from what we’ve covered, 17025 applies to labs who test and calibrate.
UL is required to adopt 17025 criteria as part of DAP as a ‘chain of custody’ application of the controls for data etc.
This adoption of 17025 requirements affects both UL and its client labs alike.

## Slide 81: Discussion and Questions
Notes: Instructor Notes:
Ask if there are any questions before moving on to the next module.

## Slide 82: Thank You. Stay SAFE!
For more information on this and additional topics, please visit:
www.UL.com/LMS or call
1.888.503.5536