## Slide 1: CAR 163915755 Observation

## Slide 2: CAR 16915755
Requirement: 5.4.7.2.a When computers or automated equipment are used for the acquisition, processing, recording, reporting, storage or retrieval of test or calibration data, the laboratory shall ensure that: a) computer software developed by the user is documented in sufficient detail and is suitably validated as being adequate for use;
Non-Conformance: The Multi-Channel Temperature and Power v14 program is used as data acquisition software for Energy Efficiency Commercial Refrigerator testing. There is no evidence there has been any software validation before implementation. 2 of 5 excel datasheet calculations reviewed had no evidence of the calculations being validated.
Objective Evidence: Interview with Quality and Reliability Manager that created the program did not perform any validation of the software program. Engineering Manager, showed the handwritten calculations to show validation for datasheet calculations, these handwritten calculations were not available for all spreadsheets. Datasheets missing hand calculations were: Refrigeration and Hot Food Cabinet.
Attachments/Comments:
Standard Category: Methods and Method Validation Owner’s Org./Function: Commercial and Industrial – Laboratories
Add a note re OOT analysis milestone.

## Slide 3: CAR 163915755
2016/03/18 01:51:26 PM - Comments - Hi Mo, This CAR is coming due for a Response by the end of the day. Please submit your response with the action items and due dates, or a request for an extension by the end of the day. If this is not completed, the CAR will escalate. Let me know if we need to discuss. Jeff by Jeffrey Follett (Administrator)
2016/03/18 02:28:40 PM - Response Date Extension From 2016/03/18 To 2016/04/18 by Moriam Raji (Owner) Reason - Meeting with all stakeholders to Review the Software to verify if program computes calculations in order to proceed.
Good reminder
To Owner
31 day
extension

## Slide 4: CAR 16915755
2016/05/26 04:52:16 PM - The milestone- 'If calculations are involved, identify any standards involved and record formulas. ' due date was extended by 35 day(s) by Jeffrey Follett (Administrator)
2016/07/13 12:31:03 PM - Implementation Approved and Closed by Jeffrey Follett (Administrator) Comments - CAR is being closed due to information needed to close the final milestone is included in this second-to-last milestone. In addition, the system would not allow me to approve the implementation of the second to last milestone without first requesting and receiving approval for an extention of the last milestone since it was already past due. The best way to handle this situation appeared to be to include information needed to close the final milestone, and approve closure of the CAR.
No record of approval for 35 day Ext.
Open discussion. Last milestone has no close date.

## Slide 5: Milestones for CAR 163915755 – Open discussion regarding the CAR Admin making most of the entries into Milestones 1 and 2. (Entries for M1 shown below.)
2016/03/24 04:02:34 PM - Milestone Expectation was changed from Screen shots of Software code and verification of formulas and/or verification of software input to output using another calibrated instrument to Screen shots of Software code used in refrigeration and hot food cabinet along with title and revision level. by Jeffrey Follett (Administrator) 2016/04/27 11:46:29 AM - Milestone Expectation was changed from Screen shots of Software code used in refrigeration and hot food cabinet along with title and revision level. to Screen shots of User interface that displays title and version of Software in refrigeration and hot food cabinet data acquisition. by Jeffrey Follett (Administrator) 2016/04/29 03:14:25 PM - Implementation Objective Evidence was added by Jeffrey Follett (Administrator)

## Slide 6
Hand calculations showed software was accurate.